## Slide 1: Korean Cyclotrons in Present and Future
CHAI, Jong-Seo
Sungkyunkwan University

## Slide 2
Korean Cyclotrons
Cyclotron Development
Medical and Industrial Applications

## Slide 4: Korean cyclorons
Cyclotron History in Korea

## Slide 5
Korean Cyclotrons
1. KIRAMS ( Korea Institute of Radiation and Medical Science )
Energy
Proton 23 ~50 MeV
Deuteron 12 ~ 25 MeV
He-3 25 ~ 73 MeV
He-4 23 ~ 68 MeV
Beam Current
~ 70 uA
RI production
Fundamental Research
First beam Mar. 1985
Manufacturer Scanditronix

## Slide 7: KIRAMS-1 & KIRAMS-13 Development

## Slide 8
2000. 10 The First Beam Test of 1 MeV Cyclotron
Energy : 1 MeV
Particle : proton
Beam Current : 100 nA
KIRAMS-1 Cyclotron

## Slide 9
KIRAMS-13 Cyclotron (1)
2002. 2. Development of 13 MeV Cyclotron for the Medical Use
Beam Energy : 13 MeV
Particles : Proton
Beam Current : 10 ㎂
RF Cavity
Vacuum Chamber
Vacuum Pump
Ion Source
RF Dee
Return Yoke
Magnet

## Slide 10
KIRAMS-13 Cyclotron (2)
Beam Energy : 13 MeV
Particles : Proton
Beam Current : 70 ㎂

## Slide 15
Regional Cyclotron Center in KOREA

## Slide 16: KIRAMS-30 Development

## Slide 17: KIRAMS 30 – General Specifications
General Specifications | 
Type of Accelerated Ions | Negative Hydrogen
Extraction method | Stripper carbon foil
Beam Energy(proton) | 15 ~ 30 MeV
Beam Current(proton) | Guaranteed 300 uA
No. of Beam lines | 4
Dual beam | available

## Slide 18: Schedule of KIRAMS-30 Development
Vacuum Chamber ( shape, degree of vacuum : < 5x10-7 torr )
Multi-cusp Ion source Extraction Current, Extraction Energy
Injection System (LEBT)
Steering, Focusing Units; Buncher; Inflector
Injection efficiency (Low loss, Beam matching)
Magnet system
Return Yoke, Coil, Hill/Valley Structure B0, Rext, Isochronous & Axial Focusing
RF system RF structure, Coupling, Tuning, Pickup, AMP
frf, Dee voltage, Gain/turn, PA power
Extraction System
2 steering magnet, 2 carbon foil loader
Vacuum System
Beam Transport Line
Switching Magnet,Quadrupoles
Target System
Liquid : F-18 ; Gas : C-11
Control System
Beam
Diagnostic
System
Cooling System
BPM, FARADAY CUP
FARADAY CUP
1 DP, 1 MP
2 CRYOPUMP, 1 MP
2004
2005
2006

## Slide 19: The magnet
Central field | 1.05T
Extraction radius | 0.736 m
Pole radius | 0.81 m
Hill, Valley gap | 0.03,0.62 m
Tune (νr, νz) | 1.05-1.1, 0.75
Hill angle | 48 °
Operating current | 130.5 A
Coil turns | 22 x 16 x 2
Coil power | 11.27 kW
Dimension | H 1.44 m, Φ2.7m
Weight | 50 ton

## Slide 20: Magnet

## Slide 21: Ion Source & Injection line
-negative hydrogen ion extraction
-max. 30 keV extraction energy
-max. 10 mA beam current
-normalized beam emittance ~0.8 mm mrad
Vacuum
Box
Plasma
Extraction
Ground Lens
Cusp
Body
Buncher
Solenoid
Einzel Lens

## Slide 22: Central Region

## Slide 23: RF cavities
RF Frequency | 63.96MHz
Harmonic number | 4
Type of Cavity | Dual Cavity vertically
Number of Cavities | 4
Number of Dees | 2
Dee Angle | ~39o
Coupling type | Capacitive
Cooling Method | Water Cooling
RF Amp Power | 50kW
Dee
Liner
Stem

## Slide 25: KIRAMS-30 Cyclotron at KAERI

## Slide 26: Commissioning
2006. 5 - 2007. 3 Installation
2007. 3 - 2008. 8 Beam commissioning 650 uA @ 1.5 MeV
2009. 9 - 2009. 10 Disassembly KIRAMS-30
2010. 3 - 2012. 8 New Installation
and Stop because of many reasons
2012. 9 - 2012. 11 Restart Beam Commissioning ; 350 uA @ 30 MeV

## Slide 27: SKKUCY-9 Development

## Slide 28
General and Beam Dynamics
General
Energy 9.6 MeV
Acceleration Particle H-
Maximum Beam Current
150 uA
First Beam Aug. 2013

## Slide 29: Magnet

## Slide 31
RF system

## Slide 33
Ion source

## Slide 34: Future Plan

## Slide 35: SKKUCY 4 for NR and BNCT
High intensity D- cyclotron for neutron generator
Application : Neutron Radiography, BNCT
Extraction particle to target : proton or deuteron
Energy of the extracting beam : 4 MeV (p) 2 MeV (d)
- beam energy required
~ 4 MeV for Be-Li hybrid target
Beam current : up to 1 mA

## Slide 36: General Specification
Magnet | 
Number of sectors | 4
B0 | 0.8 T
Extraction radius | 39.7 cm
Pole radius | 45 cm
Hill/Valley gap | 4 / 52 cm
Hill angle | > 40°
Coil dimension | 7dpcⅹ10 turns
NI / coil | 28000 A-turns
Magnet dimension | Φ150, H82 cm
Power | ~ 8 kW
Parameters | Values
Particle | D-
Extraction energy | 4 MeV
Ion source | 10 mA H- multicusp
Injection energy | 30 keV
Beam intensity | 1 mA
Target | Be-Li hybrid
RF | 
Harmonic number | 4
RF frequency | 48.83 MHz
Number of dees | 2
Dee angle | 43°
Dee voltage | 50 kV
Coupling | Capacitive
Power | <50 kW

## Slide 37: Cyclotron Viewing
 Picture of the cyclotron chamber open. The estimated mass of the machine is about 9.1 ton.
 Drawings of the magnet and RF systems assembled.

## Slide 38
Vacuum
Proton Beam
Copper
Aluminum
Cooling water
Lithium
x
y
z

## Slide 39
Analysis
Software : FLUENT
Condition
  1.7 MeV, 1 mA on Li
  Angle : 30o
  Thickness : 100mm
  Flow rate : 40 L/min, 15 ℃
Result
  Maximum Temperature : 325 K (= 52 ℃)

## Slide 40: Schematic Diagram of Neutron Radiography System
Cyclotron
Proton Beam
Be-Target
Object
Scintillation
Converter
Monitor
Monitor
Video
Recorder
CCD Camera
Image
Processor
Floppy
Disk

## Slide 41: Comparison of X-Radiography and Neutron Radiography
Neutron
X-ray(120keV)

## Slide 42: Li-ion Battery Imaging by Neutron Radiography

## Slide 43: Image of Lead-Shielded Battery by Neutron Radiography

## Slide 44: Comparison of X-ray and Neutron Radiography for Electrical Drill
neutron
X-ray

## Slide 46
2. NCC (National Cancer Center )
Energy
protron 235 MeV
Isocentric Gantries
Proton therapy
Medical Research
First beam Jul. 2005
Manufacturer : IBA

## Slide 47
가속기 시스템 소형화를 통한 비용 절감과 설치 용이
1조원
300 억원 내외
700 톤
20 톤

## Slide 49
각 연구 분야 전문성 융합
가속시스템
사이클로트론
치료기
조사시뮬레이션 및 영상진단
전시스템
운영 및 의료, 교육 S/W
의학물리
장치 개발 후 실제 임상에서 사용할 수 있는
실용적인 양성자 치료기 개발
양성자 치료
시스템
방사선종양학
개발 전략
“신뢰도 높고 안정적인 치료 시스템 개발”
임상 및 의학물리 연구 사용자 요구 기반의 장치 개발을 목표로 설정

## Slide 50
7. 기술 안정화 전략
II. 양성자 치료기 개발
사이클로트론을 이용한 양성자 치료기 개발
영남권
호남권
충청권
서울특별시
‘권역별’ 양성자 치료 센터 공급 전략
지자체 참여 유도 및 정치적 전략 접근이 용이 투자 참여의 현실성 증대
중앙의 컨트롤타워 마련 각 센터의 운영상태 분석을 통한 성능강화 피드백 시스템 구축, 연구분야 중첩 방지
타 사업과의 중복성 방지
정부차원의 관심도 상승
“ 과제 실현의 신뢰도 향상 가능”
(AMC, SMC, SNUH,
Severance, NCC, CMC)
장비
시스템
유지보수
인력
턴키 방식의 패키지화
* 예상 선정지이므로 추후 변경 가능

## Slide 51
Y. Kadi etc. CERN report

## Slide 53: Thank for Your Attention