## Slide 1
Since 2013 the operation of the 100-MeV proton linear accelerator has been started, the ACCT (AC-current transformer) has been a main beam current measurement system of the 100-MeV proton linear accelerator and its beamline. ACCT (AC-current transformer) is typical non-destructive current measurement method, can measure the macro-pulse current of the charged particle beam without any disruption [1]. And then, ACCT is superior to measure the beam pulses longer than tens of sec with high sensitivity and strong voltage signal, even down to 1 mA. Recently, new ACCT electronics is introduced and tested, which have new functionality of the in-situ calibration. In this paper, new upgraded ACCT system will be introduced.
Future Plan of ACCT system
Test for New ACCT Electronics and Plan for ACCT System
at KOMAC Space Radiation Testing Beamline
Korea Multipurpose Accelerator Complex, Korea Atomic Energy Research Institute, Gyeongju, Korea
Sang-Pil Yoon*, Young-Gi Song, Han-Sung Kim, Hyeok-Jung Kwon
KNS Spring 2022, Jeju 2022.5.18~20
* spyun@kaeri.re.kr
This work has been supported by the National Research Foundation of Korea (NRF-2021M2D1A1045615).
Test of New ACCT Electronics
Introduction & Motivation
Test and calibration set-up
Current ACCT System
Full Scale Range | Factory preset
Ratio accuracy error | < 0.1% FS
Lower cutoff | < 3Hz
Upper cutoff | 1 MHz
Rise time | 350 ns (10 – 90%)
Manufacturer | Bergoz (France)
Specification
Test and calibration set-up
Polarity control
Schematics future Installation at the space radiation test facility
Main feature of New ACCT Electronics
3 selectable full scale ranges (1 / 10 / 20 mA)
3 outputs : High impedence, 50Ω and differential
Integrated calibration current source
Remote controllable by TTL
The calibrated pulse is generated by the ACCT-E-RM-3R and injected into the sensor calibration winding via the sensor CAL input connector
Current pulse amplitude : 50 % of the full scale of the selected range
( 10mA full scale => calibration current is 5 mA)
The calibrated current pulse length is defined Trigger signal length
(trigger signal length = calibration current pulse length)
The trigger signal has to be TTL level.
CAL Trigger to ACCT-E-RM-3R : 3.3V, 500 usec by waveform generator
ACCT output voltage : measured by Oscilloscope
Preset ACCT range : 10V / 20 mA
CAL current pulse amplitude : 10 mA (50% of the selected range)
Pin 1 status | Pin 2 status | Range
Low | Low | Range 1
High | High | Range 1
High | Low | Range 2
Low | High | Range 3
Range control
Calibration result old ACCT sensor head
Remote control function
Range control
Calibration polarity
Power status output
(+5V from external power supply)
Remote control by TTL
Power status monitoring
Summary & Future Plan
In-situ calibration of ACCT without disassembling and test stand
Easy adjustment to polarity by remote control
Easy adjustment to full scale range remotely (3 selectable range)
Power status monitoring of ACCT electronics
Periodic response check and calibration of ACCT system
5V, 500 usec current pulse measured
(for input : 10 mA /10V)
Current pulse adjusted by input TTL pulse length
Old ACCT has no signal degradation
ACCT consist of toroidal sensor head and electronics
ACCT is main beam current monitor at 100 MeV proton linear accelerator